HBM

Im Tiefen See 45

64293 Darmstadt

HBM Sensordatabase 
Format Specification

Version 1.2, 27.04.2008

# Table of Contents

Table of Contents	2

Introduction	3

Example	13

Database fields and their equivalents in Property-grid	14

Necessary adjustments	18

HBM Sensordatabase
Format Specification

# Introduction

The HBM Sensordatabase is available as a Microsoft Access 2002 database. All important properties of sensors are stored in a single table (SenPara). Another table (Groups) contains the groups where the various sensors are arranged. The Groups-table is used by HBM software to build corresponding TreeViews – these entries are also not relevant for the correct amplifier configuration.

In the About-table there is a Version-field which describes the current version of the database format. This document serves to specify the version 20070802.

The following sections describe all the entries in the SenPara-table:

| Table filed and data type | Description |
| --- | --- |
| Description (Text 50 characters) | Descriptive name of the sensor. |
| GenericType (Integer) | Generic sensor type. StrainGage =1 Bridge =2 LVDT =3 Voltage =4 Current =5 Frequency =6 Thermocouple =7 RTD =8 Piezo =9 Resistor =10 Ind.Bridge =11 Incremental =12 Charge =13 Poti =14 PWM =15 Time =16 SSI =30 IEPE =31 Pressure =32 CAN =100 |
| Type (Text 50 characters) | Reserved. |
| SensorType (Integer) | Specific sensor type. This field describes several characteristics (wiring, resistance, bridge type ...) as a set of enumerations. Following are the major types of sensors: full bridge =350 half bridge =351 SG fullbridge =353 SG halfbridge =354 Single SG =355 Inductive fullbridge =356 Inductive halfbridge =357 SG fullbridge 120 Ohms =362 SG fullbridge 350 Ohms =363 SG fullbridge 700 Ohms =364 SG halfbridge 120 Ohms =365 SG halfbridge 350 Ohms =366 SG halfbridge 700 Ohms =367 Single SG 4 wire 120 Ohms =368 Single SG 4 wire 350 Ohms =369 Single SG 4 wire 700 Ohms =370 Single SG 3 wire 120 Ohms =371 Single SG 3 wire 350 Ohms =372 Single SG 3 wire 700 Ohms =373 Single SG 3 wire 1000 Ohms =374 Single SG 4 wire 1000 Ohms =376 LVDT =380 Potentiometer =385 DC voltage =420 DC current =421 Thermocouple Type J =450 Thermocouple Type K =451 Thermocouple Type T =452 Thermocouple Type S =453 Thermocouple Type B =454 Thermocouple Type E =455 Thermocouple Type R =456 Thermocouple Type N =457 Resistor =475 Pt10 =500 Pt100 =501 Pt500 =502 Pt1000 =503 Frequency =520 PWM =521 Counter F1+F2 =522 PWM F1+F2 =522 PWM F1+F2 (4x) =523 Counter F1+F2 (4x) =523 Frequency F1+F2 (4x) =523 Time =524 Counter F1 =525 XPWM PWM mode =539 XPWM Pulse width mode =540 XPWM Period mode =541 Piezoelectric (active power supplied) =555 Piezoelectric (passive) =570 SSI =580 IEPE =581 Pressure (Pressure sensor) =582 Inductive counter =590 Inductive frequency =591 |
| ElectrUnit (Text 8 characters) | The electrical unit is derived from the SensorType as follows: SensorType 350 – 385 :mV/V SensorType 420 :V SensorType 421 :mA SensorType 450 - 457 :°C SensorType 475 :Ohm SensorType 500 - 503 :°C SensorType 520 - 521 :kHz SensorType 520 - 521 :kHz SensorType 522 :kIMP SensorType 555 :V SensorType 570 :nC SensorType 580 :mV SensorType 581 :V SensorType 582 :bar SensorType 1000 :% SensorType 1001-1002 :ms |
| EngUnit (Text 8 characters) | The physical unit of measured value provided by the sensor. |
| GageFactor (Single) | The strain-sensitivity of the strain gage (ratio of relative change in resistance to stretch). It is only relevant for strain gage circuits, but not for sensors. For sensors use the default value of 0. |
| BridgeFactor (Single) | Bridge factor. It is only relevant for strain gage circuits, but not for sensors. For sensors use the default value of 0. |
| SADPara (Integer) | Variable content, depending on the generic SensorType: Input circuit for torque measurement (MGCplus/ML60) 5 V input =21 100 mV input =22 Input circuit for torque measurement (MGCplus/ML460) DIRECT = 23 INDUCTIVE = 24 Circuit (e.g. Pt100 or resistance) 3 wire =25 4 wire =26 2 wire =27 5 wire =28 6 wire =29 Discharge of the charge amplifier SHORT =31 MEDIUM =32 LONG =33 |
| Excitation (Single) | Supply voltage in V. |
| ScaleType (Integer) | Scaling type. Following are the scaling types: TwoPoint =0 (P1X, P1Y, P2X, P2Y) ZeroSpan =1 (ElecZero=P1X, ElecSpan=P2X, PhysSpan=P2Y) Table =2 (Values ​​are in ScaleData1-field, separated by ~) Polynomial =4 (Values ​​are in ScaleData2- field, separated by ~) Table correction =8 Polynomial correction =16 More Polynomials =32 (Values ​​are in ScaleData3- field, separated by ~) Gain =64 (m=P1X, b=P1Y) Off =128 Internal scale =256 |
| P1X, P1Y, P2X, P2Y (Single) | Scaling points according to the ScaleType. These fields are used for scale type TwoPoint und ZeroSpan (ScaleType=0 or 1). In the case of ScaleType=2 (Table) is the first table point in P1 and the last table point in P2. |
| Range (Single) | Physical range of measured value. Or the maximum electrical output signal in mV/V, when the sensor is a kind of strain gage bridge, and was not created by Catman SDS (which you can identify easily by taking a look at the Wiring-field. Value!=0 means that QXSDB has at last written the sensor). |
| ScaleData1 (Memo) | In the case of ScaleType=2 (Table) is like following: <Number of table values (N*2)>~ <elec.value1>~<phys.value1>~ … <elec.valueN>~<phys.valueN>~ <phys.unit (e.g. %)>~ <phys.range> For example: 8~0~0~0,5~50~0,75~78~1~100~%~100 |
| ScaleData2 (Memo) | In the case of ScaleType=4 (Polynomial) is like following: <Number of polynomial coefficients (N)>~ <a(0)>~ … <a(N)>~ <phys. unit (e.g. %)> |
| ScaleData3 (Memo) | Entries for the table correction when ScaleType=0 or 1. <Number of table values (N*2)>~ <elec.value1>~<phys.value1>~ … <elec.valueN>~<phys.valueN>~ <phys.unit (e.g. %)>~<0> For example: 8~1,1~11,1~2,2~22,2~3,3~33,3~4,4~44,4~%~0 |
| ScaleData4 (Memo) | Entries for the polynomial correction when ScaleType=0 or 1. <Number of coefficients (N)>~ <a(0)>~ … <a(N-1)>~ <phys.unit (e.g. %)>~ <0> For example: 3~1,1~2,2~4,4~%~0 |
| ScaleData5 (Memo) | In the case of ScaleType=64 (Gain) is like following: <Offset>~ <Gradient> In the case of ScaleType=32 (More Polynomials): <Number of polynomials (N)>~ <Start value of polynomial 1>~ <Number of polynomial coefficients 1 (M)>~ <a(0)>~<Exponent 0 – here always 0!> <a(1)>~<Exponent 1 – here always 1!> … <a(M-1)>~<Exponent M-1 – here always M-1)> <Start value of polynomial 2>~ … <Start value of polynomial N>~ … |
| TID (Text 32 characters) | Sensor identification (PRIMARYINDEX of the table). |
| SNR (Text 50 characters) | Serial number of sensor. |
| GroupIndex (Long) | Index of the group in Groups-table, which the sensor is associated with. |
| PositionIndex (Long) | Index of the sensor position within the Groups-table. |
| Comment (Text 50 characters) | Comment to the sensor. |
| Store (Text 50 characters) | Location of the sensor or position of the measuring point. |
| CalDate (Date/Time) | Calibration date. |
| CalExpriation (Date/Time) | Date on which the calibration validity ends. |
| Descriptor (Text 50 characters) | Description file (e.g. file path). |
| ExcitationFrequency (Long) | Carrier frequency in Hz. Exceptions: DC =-1 Auto =0 |
| Resistance (Single) | Resistance (is currently used only for thermocouple, strain gage and potentiometer sensors). |
| Wiring (Integer) | Type of sensor wiring: 2 wire =2 3 wire =3 4 wire =4 5 wire =5 6 wire =6 |
| SSIClockFrequency (Long) | Chopper frequency (if data record describes an SSI sensor): 100 kHz =0 200 kHz =1 500 kHz =2 1000 kHz =3 |
| SSIResolution | Resolution (number of bits): 12 Bit =0 13 Bit =1 ….. 32 Bit =20 |
| SSICoding | Encoding: Gray =0 Dual =1 |
| Manufacturer (Text 50 characters) | Name of the sensor manufacturer. |
| ActiveSupply (Single) | Supply voltage (for sensors that require an external power supply). |
| CounterIndexEnabled (Integer) | If a counter sensor is described: Counter does not consider the NullIndex =0 Counter reset to 0 at NullIndex =1 |
| CounterIndexDivisor (Long) | Divider for the index signal. |
| ExpectedShuntValue | Expected measuring value in the activated shunt. |
| PulsesPerRound | Impulse of a rotation (characteristic of a counter sensor). |
| GlitchFilter | Width of the Glitch-filter (counter and pulse-width sensors): Off = 0 1 = 0,1 µs 2 = 1 µs 3 = 10 µs 4 = 100 µs |
| Polarity | Polarity (for counter and pulse-width sensors) |
| ActiveSupplyVoltage(Single) | Active power for sensors in V (e.g. CAN sensors). Not to be confused with power supply for bridges! |
| CalInitials (Text 8 characters) | Initials of the sensor calibrator. |
| Supply (Double) | Reserved. |

The remaining table fields describe the characteristics of generic CAN sensors and are not relevant for all other sensor types.

| Table field and data type | Description |
| --- | --- |
| CAN_ID (Long) | CAN message ID. |
| CAN_FrameFormat (Integer) | StandardFrame =0 ExtendedFrame (29Bit) =1 |
| CAN_ByteOrder (Integer) | Byte-Order: Motorola =0 Intel =1 |
| CAN_RawValueFormat (Integer) | Data format of this signal in the CAN message: 32-bit unsigned integer =0 32-bit signed integer =1 64-bit unsigned integer =2 64-bit signed integer =3 32-bit floating-point number =4 64-bit floating-point number =5 |
| CAN_SigType (Integer) | Signal type: Standardsignal =0 Mode-dependent signal =1 Mode-signal =2 |
| CAN_StartBit (Integer) | Position of the start bit of the signal in the CAN message. |
| CAN_BitCount (Integer) | Number of bits that are used for this signal. |
| CAN_ScaleFac (Single) | Scaling factor which is used to calculate the output signal from the digitally encoded signal value (data format). |
| CAN_Offset (Single) | Offset which is used to calculate the output signal from the digitally encoded signal value (data format). |
| CAN_MinVal (Single) | Expected minimum value of the calculated signal. It can be used to monitor the under-range of the measurement. |
| CAN_MaxVal (Single) | Expected maximum value of the calculated signal. It can be used to monitor the over-range of the measurement. |
| CAN_Unit (Text 8 characters) | Physical unit to be used for this signal. |
| CAN_ModeSig (Integer) | Data format of the Mode-signal. 32-bit unsigned integer=0 (No other values ​​allowed). |
| CAN_ModeByteOrder (Integer) | Byte-Order of the Mode-signal: Motorola =0 Intel =1 |
| CAN_ModeMode (Integer) | Mode-value. |
| CAN_ModeStartBit (Integer) | Position of the start bit of the Mode-signal. |
| CAN_ModeBitCount (Integer) | Number of bits that are used for the Mode-signal. |
| CAN_MaxRepTime | Maximum response time for reply. |

# Example

There are some examples of commonly used sensors in the following table. Please note that not all fields are available for every sensor type (The unfilled fields are not decisive for the respective sensor type).

|  | Examples of various sensor types |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Table field | Temperature | Thermoelement | Pressure | Force | Displacement |
| Description | Pt100 | Thermoelement Type K | P19 4bar | C2 1kN | WA 500mm |
| GenericType | 7 | 8 | 4 | 2 | 11 |
| SensorType | 501 | 451 | 420 | 350 | 356 |
| ElectrUnit | °C | °C | V | mV/V | mV/V |
| EngUnit | °C | °C | bar | kN | mm |
| SADPara | 25 | 0 | 0 | 23 | 0 |
| Excitation | 20 | 0 | 0 | 5 | 2,5 |
| ScaleType |  |  | 0 | 0 | 0 |
| P1X |  |  | 0 | 0 | 0 |
| P1Y |  |  | 0 | 0 | 0 |
| P2X |  |  | 10 | 2 | 80 |
| P2Y |  |  | 4 | 1 | 500 |
| Range |  |  | 4 | 1 | 500 |
| TID | Pt100 | Thermo_K | HBM_P19 _4bar/a | HBM_C2 _1kN | HBM_WA _500mm |
| GageFactor | 0 | 0 | 0 | 0 | 0 |
| BridgeFactor | 1 | 1 | 1 | 1 | 1 |
| GroupIndex | 70 | 70 | 47 | 12 | 41 |
| PositionIndex | 10 | 2 | 1 | 3 | 16 |

# Database fields and their equivalents in Property-grid

The following table lists the equivalents of the database fields to the row labels in the Property-grid.

Some database fields match none of the fields in the Property-grid, some have variable contents and are dependent on the sensor type.

| Name of Database field | Description in Property-grid |
| --- | --- |
| Description | Description (in category Identification) |
| GenericType | 1 = SG bridge sensor 2 = Bridge sensor 3 = LVDT sensor 4 = Voltage sensor 5 = Current sensor 6 = Frequency sensor 7 = Pt sensor 8 = Thermoelement sensor 9 = NO-EQUIVALENT 10 = Resistance sensor 11 = Inductive bridge sensor 12 = Counter sensor 13 = NO-EQUIVALENT 14 = Potentiometer sensor 15 = NO-EQUIVALENT 16 = NO-EQUIVALENT 17 = NO-EQUIVALENT 30 = SSI sensor 31 = IEPE 32 = Pressure sensor 33 = Inductive counter sensor 34 = Inductive frequency sensor 100 = CAN sensor |
| Type | Model (in category Identification) |
| SensorType | Bridge type, bridge circuit type, resistance in Ohms, thermocouple type. Depending on the sensor ... |
| Electrical Unit | Electrical unit. |
| EngUnit | Physical unit. |
| GageFactor | k-factor (by the bridge sensor). |
| BridgeFactor | Bridge factor (by the bridge sensor). |
| SADPara | Circuit type ... |
| Excitation | Supply voltage in V. |
| ScaleType | Active scaling type (in category Scaling). |
| P1X | "Electrical value" (e.g. when active ScaleType=Table). |
| P1Y | Physical value. |
| P2X, P2Y | According to P1X, P1Y. |
| Range, Scaledata1...Scaledata5 | No direct equivalent. |
| TID | Sensor-ID |
| SNR | Serial number. |
| GroupIndex, PositionIndex | No direct equivalent. |
| Comment | Comment (in category Identification). |
| Store | Measuring point identification. |
| CalDate | Calibration date. |
| CalExpiration | Calibration period (CAUTION: Internally is here the date up to which the calibration is valid ...). |
| Descriptor | Description file. |
| CAN_ID | CAN-ID. |
| CAN_FrameFormat | Frame format. |
| CAN_ByteOrder | Byte-Order. |
| CAN_ByteCount | No direct equivalent. |
| CAN_RawvALUE | Data format. |
| CAN_SigType | Signal type. |
| CAN_StartBit | Start bit. |
| CAN_BitCount | Number of bits. |
| CAN_ScaleFactor | Scaling factor. |
| CAN_Offset | Scaling offset. |
| CAN_MinVal | Minimum value. |
| CAN_MaxVal | Maximum value. |
| CAN_Unit | Physical unit or custom output unit. |
| CAN_ModeSig | Mode-Data format. |
| CAN_ModeByteOrder | Mode-Byte-Order |
| CAN_ModeMode | Mode-Value. |
| CAN_ModeStartBit | Mode-Start bit. |
| CAN_ModeBitCount | Number of bits for Mode-signal. |
| ExcitationFrequency | Carrier frequency (e.g. by SG bridge sensor). |
| Resistance | Resistance in Ohm. |
| Wiring | Wiring type. |
| SSIClockFrequency | Chopper frequency in Hz (by SSI sensor). |
| SSIResolution | Resolution (by SSI sensor). |
| SSICoding | Encoding (by SSI sensor). |
| Manufacturer | Manufacturer. |
| ActiveSupply | Supply voltage. |
| CounterIndexEnabled | Signal for NullIndex. |
| CounterIndexDivisor | Divider for the index signal. |
| CalIntials | Initials of the calibrator. |

# Necessary adjustments

The Sensordatabase (the Access file) is used by the Sensordatabase-engine of the QuantumX Assistant, as well as by Catman. The clients do not use the same code. It is therefore important, changes that are made to the actual database (the Access file) must be considered in both clients.

QX DB-Client:

The following fields are not yet considered:

- ExpectedShuntValue

- GlitchFilter

- Polarity

- CAN_MaxRepTime

To remain compatible with the Catman DB client, the fields exist however. The fields are added when converting older databases. The fields will not be overwritten when you change sensors.

Catman DB-Client:

ScaleData1...Scaledate5 are Textboxes with maximum 255 characters. This causes problems with some scaling types (e.g. table), since this maximum size can easily be exceeded.

Workaround: Convert ScaleData-fields to Memo-fields! (QXDBClient converts these fields automatically in Memo-fields).